## Materials
OFRR Ventilator Bill of Materials
Part Name | Description, manufacturer | Supplier | MFG part number | Quantity | Cost for Required quantity | Product/site info | Notes
Sea-Long
Multi-patient helmet non-latex with port | Adult-size helmet | Sea-Long.com | PN5300 | 1 | 147.98 | Sea-Long
Amron
Oxygen Treatment Hood | Adult-size helmet | Amron | 8891-01, or-02, or -03 | 1 | 150 | Amron
Bacteria/viral/noise fliters
HEPA filters | Gibeck Iso-Guard HEPA, 99.9999%, 1.6-2.3 cm H20 resistance | Teleflex, Kemala Thompson-Knot, 206 735 9690 or home (206) 829-9392 , cs@teleflex.com, | various |  | Availability is mixed
Below are common filters with lower protection, likely allow more virus through
| Sea-Long | PN154 | 10 | 32.5 | Sea-Long | 20 per case
| ViroMax, 99.99% virus, 1.5 cm H20, 22mm | A-M Systems, Sequim WA! 360 683-8300 | 156200 | 50 | 73 |  | Cases of 20- 50
| AirLife disposable bacteria filter | Vitality Vedical | 1851 | 40 | 71.46 |  | $5 for 1-4
| About 1.5 cm H20 resistance | Teleflex, NC, 866-246-6990, cs@teleflex.com | 1605 (22mm), 1603 (30 mm, 1577 (pivots, 22 mm) |  | Must inquire |  | 99.9999 viral
|  | DirectHome Medical | H1605, HCPS code A7038, GTIN/UPC code 04026704399208 | 50+ | 1.95 | Manufacturer Hudson RCI, Malaysia | 99.9999 viral
| Medtronic DAR mechanical Large | https://www.medtronic.com/content/dam/covidien/library/us/en/product/ventilator-filters/dar-sales-sheet.pdf | 351U5410 |  |  |  | 99.9999 viral
| Medtronic DAR - Mechanical Compact | https://www.medtronic.com/content/dam/covidien/library/us/en/product/ventilator-filters/dar-sales-sheet.pdf | 351U5878 |  |  |  | 99.999 viral, see sheet for resistance
| Medtronic DAR Mechanical Small | https://www.medtronic.com/content/dam/covidien/library/us/en/product/ventilator-filters/dar-sales-sheet.pdf | 351U5979 |  |  |  | "recommended for SARS" , 2cm at 60lmin, max life 48 hrs
| Medtronic DAR Electrostatic Large | https://www.medtronic.com/content/dam/covidien/library/us/en/product/ventilator-filters/dar-sales-sheet.pdf |  |  |  |  | 2cm at 60l/min, 48 hr max life
| Pall Ultipor 100 Extended Life | https://shop.pall.com/us/en/medical/mechanical-ventilation/ultipor-100-zidgri78lig | BB100A | 50/case
| Pall Bi-directional | https://shop.pall.com/us/en/medical/mechanical-ventilation/pall-bb50t-breathing-circuit-filter-zidgri78lyp | BB50T | 50/case
PEEP valve, adjustable |  | CareFusion | 2K8082 | 1 | 8.39 | CareFusion | Make sure you get the 22 mm version
Maintain pressure in helmet - 0-20cm | Ambu Inc. | Cascade Healthcaresolutions | Manu no: 199003020 | 20 | 134 |  | Make sure you get the 22 mm version
| Ambu Inc. | Buyemp Emergency Medical Products | 199-003-020 | "case" | 149.99 |  | Make sure you get the 22 mm version
| Teleflex, 5-20 cm H20 (pg 70 of catalog) | Teleflex | 5383 22mm | 10 |  |  | Make sure you get the 22 mm version
Alternative suppliers include Amazon
Corrugated tubing | Sea-Long part no. | Sea-Long | PN148 | 1 |  | Sea-Long
| Medline | Cascade Healthcaresolutions |  | 100 ft | 19.62
| Medline | Cascade Healthcaresolutions |  | 6 ft, case of 50 | 45.68 |  | Alternative supplier
Oxygen compatible lubricant | Sea-Long part no. | Sea-Long 2 oz | P1000 | 1 | 30 | Sea-Long
For o-rings prior to assembly | Multiple size tubes available | Dive Gear | ASL-PC71-97xx | 1 | $8-35 | Dive Gear O2-compatible lubricant
Optional Parts
Pressure Line Adapter, sice port to inject O2 |  | Sea-Long | PN 1031 | 1 | 1.75
Replaceable non-latex neckseal |  | Sea-Long | PN512 | 1 | 44.95
Tape-on neckdam sleeve to reduce leaks |  | Sea-Long | PN597 | 1 | 75.95
Oxygen tubing & Cannula |  | Teleflex, NC, 866-246-6990, cs@teleflex.com
Bubble humidifier |  | Teleflex, NC, 866-246-6990, cs@teleflex.com
One way valve | Sea-Long part no. | Sea-Long | PN175 | 1 | 3.25 | Sea-Long
One-way entry port protection against backflow | Tleflex Medical | Cascade HealthSolutions | 1665EA | 1 | 1.22

## PEEP of HEPA filters
Note that filters add back-pressure, which can become higher in time with accumulation of particles, or water. If the interior of the helmet has liquid water, this can suddenly block the filter, leading to high PEEP, and lack of air movement.
Back Pressure when New
| l/min flow | Mech Large | Mech Compact | Mech Small | Electro Large | Elec Small
| 30 | 0.8 | 0.8 | 1.2 | 0.6 | 0.8
| 60 | 2 | 1.9 | 2.7 | 1.5 | 2.1
| 90 | 3.6 | 3.2 | 4.5 | 2.6 | 3.7
estimated | 120 | 5.7333 | 4.7364 | 6.66 | 3.9444 | 5.6692

## Version control information
OFRR Bill of Materials
Document updated/edited by: John Selker
Most recent update: 13 April 2020